## Slide 1: Distributor training MDR Article 14

## Slide 2: Distributor Training: MDR Article 14 Compliance
Ensuring Alignment with the Legal Manufacturer’s Obligations
Presented by: Hildur Annika Jacobsen & Dr. Philip James Davenport
Date: 2025.03.31
Notes: Welcome to this training presentation on the requirements for distributors that are contained within the European Medical Device Regulations. In particular, we will explain how to comply with MDR Article 14 that specifically covers distributors, and how to work effectively with us here at SP Medical. My name is Dr Philip Davenport, and I’m a regulatory specialist, and I’m here with Hildur Jacobsen, the quality and regulatory co-ordinator with SP Medical.

## Slide 3: Agenda
Introduction to EU MDR and Article 14
Distributor Responsibilities under MDR
Required Procedures
Distributor Vigilance Obligation
Best Practices for Distributors
Common Errors
Notes: To begin, I will talk about the legal requirements which come from the European regulation on medical devices and which has particular responsibilities for those who are classed as distributors in article 14 of the MDR.
I’ll hand over to Hildur, who will cover the procedures that we think are necessary to implement to meet those requirements and what we at SP Medical expect in terms of collaboration and assistance, particularly around device vigilance and what we see as best practice in this area.
To finish, I’ll talk a little about common errors and issues that we see and possible ways to address them.

## Slide 4: Introduction to EU MDR and Article 14
What is MDR?
Regulation (EU) 2017/745 on Medical Devices.
Focus: ensuring safety, performance, and compliance throughout the supply chain.
What is Article 14?
Defines general obligations for distributors in the EU market.
Aims to ensure compliance, safety, and traceability of medical devices.
Definition
Article 2 (34) ‘distributor’ means any natural or legal person in the supply chain, other than the manufacturer or the importer, that makes a device available on the market, up until the point of putting into service.
Notes: The regulation that covers medical devices on the European Union market is the Medical Device Regulation 2017 745. This was introduced several years ago now to address a range of issues in the medical device industry, and means that there are more stringent requirements for many products, activities and economic operators on the European market. One of these was to increase the surveillance and control over all aspects of the supply chain, from the manufacturer to importer to distributor to user. Of these, Article 14, covering distributors, is the main focus of this presentation.
Article 14 contains specific, legal requirements for those classed as distributors – that is any actor in the supply chain that is not classified as a manufacturer or importer and that makes a medical device available on the European market, right up until the point that the device is actually placed into use.
Although the MDR is a long and often complex document, the requirements for distributors are mostly common sense and straightforward. The purpose of article 14 is to ensure that whenever devices are under the control or the influence of a distributor that they stay in compliance with the other requirements of the MDR, that products remain safe to use for patients and clinicians and that devices are able to be traced throughout the supply chain so that in the event of an issue, effective action can taken to resolve the problem.

## Slide 5: Distributor Responsibilities under MDR and Article 14
Key Obligations:
Verification of Compliance (Article 14.2)
Storage and Transport (Article 14.3)
Handling non-conforming Devices (Article 14.4)
Complaint Handling (Article 14.5)
Information Sharing (Article 14.6)
Notes: Article 14 has several clauses that describe specific obligations for distributors. Generally speaking, these cover
Firstly, verifying that the product and its documentation meets legal requirements
Storing and transporting products in appropriate conditions
How to handle any non-conforming products
How to deal with complaints
And lastly, when and how to share information with other parties
I’ll go through each of these clauses in turn before Hildur covers what processes we expect to be in place to manage this. Firstly, I’ll go through verification of compliance.

## Slide 6: Verification of Compliance
1. Verification of Compliance (Article 14.2)
CE Marking: Confirm that the device bears a valid CE mark. This is typically printed on the device, its packaging, or documentation.
EU Declaration of Conformity: Request and keep a copy of the EU Declaration of Conformity for each product. This document confirms the manufacturer’s compliance with MDR.
Labeling and Instructions: Verify that the device label and instructions for use (IFU) are in the required language(s) for the destination country and include all mandatory information.
Importer Information (when applicable): Ensure that the importer’s name, address, and contact details are visible on the packaging or documentation.
It is permissible to use a sampling method that is representative of the devices that are supplied
Notes: So, in Clause 14.2, Distributors have to confirm a range of documentation is present for the device. In particular:
That the device has a valid CE mark on the packaging, labelling or on the device itself, and that this matches the rest of the documentation.
Each product must have a document, the declaration of conformity. This is produced by the manufacturer and contains a range of information on the device and how it meets the requirements of the medical device regulations. You should have an up-to-date version of this document readily accessible for all devices you distribute, stored in an organized manner so that information can be found quickly whenever it is needed.
The contents of the labelling and instructions for use have a required set of mandatory information. On top of this, some countries may have particular requirements of their own– for example, for the languages that the material is available in. You should confirm that the material supplied with the device meets all the requirements for the countries that you supply. You should also be able to recognize significant omissions or other labelling issues.
Whenever applicable, distributors are also obliged to ensure that importer details are present on packaging and documentation.
The MDR makes it clear that distributors are permitted to use a sampling method to check that these requirements are met – so you don’t need to check every single shipment, but you should have a process that routinely confirms the documentation requirements are being met.

## Slide 7: Verification of Compliance (Article 14.2)
1. Verification of Compliance
If you have cause to believe that a device is not in compliance, you should not distribute it further.
You should inform the manufacturer.
If you consider the device to present a serious risk, or is a falsified device, you must inform the national competent authority of the country where you are established.
Notes: As part of this verification of compliance, the MDR states that distributors are obliged to not distribute any further devices you think are out of compliance. In these cases, you should get in touch with the manufacturer to discuss the issue further. This is intended to protect those further down the supply chain from devices that may pose risks to the users or patients.
If you believe that the device presents a serious risk to patients or users, or if you believe that you are in possession of a falsified device, then you additionally need to inform the national competent authority for medical devices in the country where you are established.

## Slide 8: Storage and Transportation Conditions (Article 14.3)
2. Storage and Transportation Conditions
Preserve Product Integrity: Devices must be stored and transported in conditions that comply with the manufacturer’s specifications (e.g., temperature, humidity, light exposure) while it is under your responsibility.
Notes: Appropriate storage and transportation of devices can be critical to preserve product integrity and to ensure that they deliver the expected safety and performance for the user.
Distributors have to store and transport products in accordance with what is specified by the manufacturer – this can include a temperature and humidity range, and to be kept out of sunlight for example.

## Slide 9: Non-Conforming Device (Article 14.4)
3. Non-conforming Device
If you suspect a device is non-conforming, you need to inform the manufacturer.
You must co-operate with the manufacturer and any competent authority in any necessary corrective actions to bring products into conformity, withdrawal or recall.
Notes: According to MDR Article 14.4, you must inform the manufacturer if, for whatever reason, you come to believe a device doesn’t conform with the requirements of the MDR. This is obviously the first step in resolving the issue and finding a solution.
Distributors must additionally co-operate with the manufacturer and applicable national competent authorities in the event of a corrective action being implemented – whether this is in some rework of the device, correcting information supplied with the device, withdrawal of the product or a recall from the supply chain.

## Slide 10: Complaint Handling (Article 14.5)
4. Complaint Handling
If you receive complaints or reports from healthcare professionals, users, etc., about a suspected incident related to a device you distribute, you must immediately forward this information to the manufacturer.
You must keep a register of complaints, any non-conforming devices, recalls and withdrawals.
You must keep the manufacturer informed regarding this monitoring and supply any information regarding this on request.
Notes: Distributors also have a role in handling complaints information. If you receive complaints about possible incidents from your customers, users or healthcare professionals, these need to be forwarded to the manufacturer without delay.
The MDR also obliges distributors to keep a register of complaints received, along with records of any of the non-conforming devices, withdrawals or recalls that I mentioned previously.
You also have a requirement to keep the manufacturer up-to-date regarding these actions, and supply the information you hold on these situations on request to support our own processes.

## Slide 11: Information sharing (Article 14.6)
5. Information Sharing
You should supply any information and documentation in your possession that is necessary to demonstrate conformity of a device to a national competent authority if it is requested.
This can be fulfilled with the assistance of the manufacturer.
You must cooperate with national competent authorities to eliminate risks posed by devices that you distribute.
National competent authorities may request access to samples.
Notes: On the topic of information sharing, the national competent authorities are also entitled to ask for certain information that you may possess.
This may be to confirm that a particular product is in conformity (using the documentation described in relation to clause 14.2), or it may be related to information on a field corrective action. As the legal manufacturer, we can help you answer these queries in order to meet the requirement.
Cooperation is necessary from the point of view of the national competent authority to eliminate risks across the supply chain. They may need to be aware of the position and condition of devices that distributors have been involved in, and so the MDR makes that cooperation a requirement.
National competent authorities are also entitled, through Article 14 Clause 6, to have access to product samples for their own investigations if they request them.

## Slide 12: Collaboration with the Legal Manufacturer
How Distributors Support Compliance:
Post-Market Surveillance:
  Sharing information about device performance and safety.
Traceability:
  Maintaining records and ensuring device history is documented.
Incident Handling:
  Immediate reporting of serious incidents or safety concerns.
Notes: Article 14 of the MDR therefore requires distributors to take internal action, but also requires distributors to work closely with the legal manufacturer. To support the manufacturers compliance with the MDR, distributors have to contribute to post-market surveillance by sharing information that they receive regarding device performance.
They must also support the manufacturer by collecting and maintaining records that enable devices on the market to be effectively traced to the end user, so that field safety corrective actions can be managed appropriately.
Distributors also need to work with the manufacturer in reporting safety issues. These need to be reported to the manufacturer without delay so that the manufacturer can respond effectively and meet the reporting timelines.
I’ll hand over to Hildur, who will talk specifically about required procedures and vigilance obligations.

## Slide 13: Required Procedures
Under MDR (EU) 2017/745, medical device distributors are required to establish and maintain several procedures to ensure compliance with the regulation. Below is a list of required procedures with practical details:
1. Verification of Device Compliance
Distributors must establish a procedure to verify the compliance of devices before making them available on the market.
Key Elements:
Check for CE marking, a valid Declaration of Conformity and a Unique Device Identifier (UDI).
Verify that labeling and instructions for use (IFU) comply with MDR requirements and are in the local language(s).
Ensure the importer’s information is included (if applicable).

## Slide 14: Required Procedures
2. Storage and Transportation
Distributors must establish procedures to ensure devices are stored and transported according to the manufacturer’s specifications.
Key Elements:
Maintain environmental controls (i.e. temperature, humidity) when this is required by the manufacturer.
Prevent physical damage or contamination during handling and storage.
3. Complaint Handling
A procedure must be in place to manage complaints about devices and communicate them to the manufacturer and other economic operators.
Key Elements:
Log all complaints received, including product details, nature of the complaint, and actions taken.
Forward complaints incl. affected product(s) to the manufacturer as required.

## Slide 15: Required Procedures
4. Handling of Non-Conforming Devices
Distributors must have a procedure for dealing with devices that are suspected to be non-compliant.
Key Elements:
Quarantine non-conforming devices to prevent distribution.
Notify the manufacturer, authorized representative, importer, and competent authorities if needed.
Support corrective actions, such as recalls.
5. Recall and Field Safety Corrective Actions (FSCAs)
Procedures must outline how to handle and support recalls and corrective actions initiated by the manufacturer or authorities.
Key Elements:
Notify customers of recalls promptly.
Track and document the return or removal of recalled devices.
Cooperate with the manufacturer and competent authorities during the process.

## Slide 16: Required Procedures
6. Traceability and Record Keeping
Distributors must establish procedures to ensure traceability of devices and maintain records.
Key Elements:
Record all economic operators in the supply chain (manufacturer, importer, other distributors).
Maintain records of devices supplied (including batch/serial numbers, where applicable) for at least 10 years (15 years for implantable devices).
7. Post-Market Surveillance (PMS)
Distributors must contribute to the manufacturer’s post-market surveillance by forwarding feedback or reports of adverse events.
Key Elements:
Establish a feedback loop for device performance and safety.
Forward adverse event reports and observations to the manufacturer or importer.

## Slide 17: Required Procedures
8. Training and Competence Management
Distributors must ensure that staff involved in distribution are trained on MDR requirements.
Key Elements:
Document training sessions and attendance.
Regularly update staff on regulatory changes and company procedures.
9. Communication with Competent Authorities
Procedures must cover how to cooperate with authorities during audits, inspections, or investigations.
Key Elements:
Provide requested documentation promptly (e.g., traceability, compliance records).
Support incident investigations.

## Slide 18: Required Procedures
10. UDI Management
Procedures should ensure that devices requiring Unique Device Identification (UDI) are properly documented and that UDI is included on the labels and in records.
Key Elements:
Verify the presence of UDI on devices.
Record UDI details in inventory systems.
11. Periodic Procedure Review and Updates
Distributors must review their procedures periodically to ensure compliance with MDR and adjust for changes in regulations or business practices.
Key Elements:
Conduct regular internal audits.
Update procedures based on feedback, regulatory updates, or non-conformance findings.

## Slide 19: Required Procedures
12. Internal Audit
Distributors must ensure proper annual inspection of implementation of requested procedures through internal audit programme. Audit results shall be properdocumented.
13. Document Control
Distributors must ensure unique document and records control, numbering, approvals and archiving.
14. Vigilance Reporting to Local Competent Authority
Distributors, in their country of operation, must aid SP Medical in ensuring timely reporting of serious incidents to the local Competent Authority.
By implementing these procedures, distributors can systematically ensure compliance with MDR, minimize risks, and demonstrate accountability during audits and inspections.

## Slide 20: Distributor Vigilance Obligation
Under MDR (EU) 2017/745, distributors play a critical role in the vigilance system, ensuring that any incidents related to medical devices are properly reported. Here’s an overview of vigilance reporting requirements for distributors, including their obligations, key steps, and practical guidance:
1. Distributor's Vigilance Obligations
Distributors are not directly responsible for submitting vigilance reports to Competent Authorities (this is the manufacturer’s responsibility). However, distributors must:
  Identify and report serious incidents and safety concerns related to medical devices to the manufacturer and, if applicable, the importer.
  Cooperate with the manufacturer and Competent Authorities during investigations.
  Keep records of incidents and related communications for traceability.
The distributor serves as SP Medical's local representative and has a crucial role to play in serious incident reporting. Submitted reports may need to be in the local language, and so the distributor must: acknowledge their responsibilites, have a defined vigilance reporting procedure, and ensure collaboration with SP Medical.
All reporting must be coordinated closely with SP Medical to maintain compliance and consistency in communication with authorities.

## Slide 21: Distributor Vigilance Obligation
2. What Incidents Must Be Reported?
Distributors must report the following to the manufacturer:
Serious Incidents: Any event that has led to, or could have led to:
  Death.
  Serious deterioration in health.
  Serious public health risk.
Safety Concerns: Adverse effects, malfunctions, or improper use indicating a potential issue with the device.
3. Reporting Process
Step 1: Recognize an Incident
Collect details from healthcare professionals, patients, or users about the incident.
Ensure that staff can differentiate between:
  A serious incident requiring immediate attention.
  A general complaint about device performance.

## Slide 22: Distributor Vigilance Obligation
Step 2: Record the Incident
Document the following details in an Incident Report Form:
  Device details (name, UDI, batch/serial number).
  Incident description (what happened, when, where).
  User/patient details (without compromising privacy).
  Distributor actions taken (e.g., device isolation, communication).
Step 3: Report to Manufacturer/Importer
Notify the manufacturer and importer without delay upon learning of a serious incident.
Provide all documented details to enable the manufacturer to assess and report the incident to the Competent Authority.
Step 4: Support Investigations
If requested, assist the manufacturer and Competent Authorities by:
  Providing additional details about the incident.
  Supplying documentation (e.g., traceability records).
  Retrieving affected devices for further analysis.

## Slide 23: Distributor Vigilance Obligation
4. Timelines for Reporting
While the distributor's timeline to inform the manufacturer is not explicitly specified in MDR, incidents should be communicated immediately upon discovery to avoid delays in the manufacturer’s reporting obligations:
Serious Public Health Risk: Manufacturer must notify within 2 days.
Death or Serious Deterioration: Manufacturer must notify within 10 days.
Other Serious Incidents: Manufacturer must notify within 15 days.
Distributors must act promptly to allow the manufacturer to meet these deadlines.

## Slide 24: Distributor Vigilance Obligation
5. Records to Maintain
Distributors must maintain a system for documenting and tracking incidents, including:
Incident reports submitted to manufacturers/importers.
Follow-up actions taken by the distributor.
Responses or corrective actions initiated by the manufacturer.
Retention Period:
At least 10 years after the last device was made available (or 15 years for implantable devices).

## Slide 25: Distributor Vigilance Obligation
6. Training for Staff
Ensure staff involved in vigilance reporting:
Can recognize serious incidents and safety concerns.
Understand the reporting process and required timelines.
Are familiar with MDR definitions and their implications.

## Slide 26: Distributor Vigilance Obligation
7. Common Errors to Avoid
Delays in Reporting: Waiting too long to inform manufacturers of incidents.
Inadequate Documentation: Failing to collect sufficient information about the incident.
Misclassification of Complaints: Treating serious incidents as routine complaints.
Poor Communication: Not following up with manufacturers or authorities when needed.

## Slide 27: Distributor Vigilance Obligation
8. Example Incident Workflow
A distributor receives a report from a hospital about a patient injury linked to a device malfunction.
The distributor documents the issue, including:
  Device details (UDI, batch/serial number).
  Description of the incident and consequences.
The distributor isolates any remaining stock of the device (if applicable).
The incident is reported to the manufacturer and importer immediately.
The distributor supports the manufacturer in their submission to the Competent Authority and any subsequent investigation.

## Slide 28: Distributor Vigilance Obligation
9. Conclusion
While the manufacturer is the primary party responsible for submitting vigilance reports to authorities, distributors play a key role by acting as the first point of contact in the supply chain. Timely and accurate reporting, combined with robust internal procedures, ensures compliance with MDR vigilance requirements and contributes to patient safety.

## Slide 29: Common Errors
Here are the most common errors distributors make when trying to comply with MDR Article 14 and how to mitigate them:
1. Inadequate Verification
Error: Failing to ensure that labeling and instructions for use (IFU) are in the correct language or comply with MDR requirements (e.g., missing UDI, warnings, or manufacturer/importer details).
Why it happens: Lack of training or reliance on suppliers to handle labeling requirements.
Solution:
  Develop a pre-distribution checklist to verify that all required elements are present, including CE marking, UDI, and correct language.
  Train staff to spot possible errors, such as missing symbols or incorrect translations.
  Keep verified compliance documentation readily accessible, in case this is requested by an authority.
  Communicate with the manufacturer.
Notes: To close out, we’ve put together some commonly experienced issues that distributors may encounter when complying with the requirements of the MDR, as well as possible solutions. Hopefully, this will illustrate why the procedures and requirements we’ve discussed in this presentation are necessary.
Firstly, inadequate verification of the device labelling. The issue arises when there is a failure to ensure that the labelling or informational material for the device is correct (for example if it is not available in a required language). This can happen due to a lack of awareness of the applicable requirements, or assuming a supplier has met all the local laws.
One solution is to create and use a checklist that verifies all required aspects of the labelling, and training staff to spot errors that can occur. Keep this verification documentation in an accessible location so this can be retrieved and communicated to those who request it.

## Slide 30: Common Errors
2. Overlooking Importer Information
Error: Distributing devices without confirming the importer’s name and address are present on the device or accompanying documentation.
Why it happens: Assuming the information is pre-checked by the manufacturer or importer.
Solution:
  Implement a mandatory review process for each shipment.
  Refuse to accept devices that lack required information.
3. Failing to Identify Non-Conforming Devices
Error: Distributing devices that do not comply with MDR (e.g., missing CE mark, incomplete Declaration of Conformity, or expired certificates).
Why it happens: Limited knowledge of compliance requirements or lack of a structured inspection process.
Solution:
  Conduct regular staff training on MDR requirements.
  Use a documented inspection checklist to verify device conformity before distribution.
Notes: For devices which additionally require importer information, this can be easily missed by distributors. It is easy to assume that this will have been checked by someone else. If importer information is required, then it may be most easily checked when the shipment is received, and a standard process or form may facilitate this. Rejecting devices that don’t contain this information may also be necessary.
Another error would be to distribute devices that are non-compliant with the MDR on the European market. This may be due to a limited amount of knowledge of these requirements around the necessary documentation, or an inadequacy of the incoming inspection process if this detail is missed. Therefore, the solution may be to have that documented inspection process, using checklists were appropriate. Ensuring that staff are appropriately aware of the requirements of the MDR is important to be sure that the products being supplied are suitably documented.

## Slide 31: Common Errors
4. Poor Complaint Handling
Error: Not maintaining an adequate system for capturing and forwarding complaints the manufacturer.
Why it happens: Lack of clarity on reporting responsibilities or absence of a structured complaint system.
Solution:
  Implement a centralized complaint log or complaint management software to document and track issues.
  Assign a designated complaint coordinator to ensure timely reporting to manufacturers and competent authorities.
5. Inadequate Record Keeping
Error: Failing to maintain records of devices supplied or received, or being unable to trace economic operators in the supply chain.
Why it happens: Reliance on manual processes or incomplete documentation systems.
Solution:
  Use inventory management software to automate record-keeping and ensure traceability. Integrate UDI where possible.
  Maintain an up-to-date customer database for recall communications.
  Periodically audit records to ensure accuracy and completeness.
  Simulate recall scenarios to test the effectiveness of your procedures.
Notes: Failing to handle complaints effectively is another common issue, particularly in the capturing and forwarding of complaints to the manufacturer. This is generally because of a lack of a structured system to properly collect and manage this information. Establishing a robust, centralized log to document these issues is recommended, and having a clear line of communication and responsibility for this communication will help make sure that the issues are handled effectively.
Some organizations have difficulty with adequate record keeping and are unable to effectively trace particular products or where they have been supplied to. Again, this is often attributable to incomplete processes and can be mitigated by developing and confirming these systems. Auditing records on a periodic basis will also confirm that these processes are working. Something that has been useful in our own system has been to test it with a simulated recall to see whether we can identify affected products and communicate the right information with the right stakeholders at the right time.

## Slide 32: Common Errors
6. Non-Compliant Storage and Transport Conditions
Error: Failing to store or transport devices according to the manufacturer’s specifications (e.g., improper temperature or humidity levels).
Why it happens: Lack of monitoring systems or misunderstanding of storage requirements.
Solution:
  Install environmental monitoring systems (e.g., temperature and humidity loggers).
  Develop standard operating procedures for product handling, including addressing any deviations
  Train staff on the specific storage and transport requirements for each device.
7. Ignoring PMS Feedback
Error: Not reporting adverse events or user feedback to the manufacturer or assuming it is not their responsibility.
Why it happens: Misunderstanding the distributor’s role in post-market surveillance.
Solution:
  Create a clear procedure for handling and reporting feedback to manufacturers.
  Use a feedback form for users to report issues.
  Communicate regularly with manufacturers about their post-market surveillance requirements.
Notes: Another error is in failing to comply with manufacturers requirements for storage and transportation of the device. This may be because the requirements are misunderstood, or perhaps simply because these values are not measured. This could be resolved by installing and using these systems developing processes to understand and monitor these conditions, and training staff on the requirements for these devices.
Distributors may also fail to fully appreciate their responsibilities regarding postmarket surveillance and not appropriately report any complaints, user feedback or adverse events back up the chain to the manufacturer. This can be due to a lack of awareness of these requirements. As with many of these issues, developing a clear procedure that describes what information is important and how to handle this communication is key, as is being open to discussing any issues with the manufacturer.

## Slide 33: Common Errors
8. Relabeling or Repackaging without Authorization
Error: Altering labels or repackaging devices without following MDR processes or getting the manufacturer’s authorization.
Why it happens: Efforts to make products market-ready faster or adapting labeling to local requirements without proper consultation.
Solution:
  Establish a strict policy against unauthorized changes to labels or packaging.
  Work closely with manufacturers to approve necessary changes in compliance with MDR.
9. Failing to Quarantine Non-Compliant Devices
Error: Continuing to distribute devices that are suspected to be non-compliant instead of quarantining them.
Why it happens: Pressure to meet delivery deadlines or lack of procedures for isolating devices.
Solution:
  Create a quarantine protocol for non-compliant devices.
  Use a non-conformance log to track actions and communication.
  Assign a compliance officer to oversee investigations and ensure no non-compliant products are distributed.
Notes: A serious error is to relabel or repackage products without authorization from the manufacturer, as this can lead to non-compliance with the medical device regulation. This can be done for several reasons, mostly around convenience or to quickly address local requirements. Distributors should have a strict policy that prevents unauthorized changes, and any changes should be made with close consultation with the legal manufacturer to preserve compliance of the device.
Failing to do this means that you may risk being considered as a device manufacturer in your own right (with all the obligations that come along with this) or coming under the additional requirements described in MDR Article 16 – these include developing a sufficient quality management system, obtaining a certificate from a notified body and a whole range of other quite onerous tasks, so it is worthwhile making sure.
Issues may also arise if distributors fail to quarantine devices effectively and they continue to reach customers. This can be because of a lack of a quarantine protocol, or other pressure from customers and can cause many issues for users and patients. Having a clear protocol and space to separate affected devices will act to reduce this risk.

## Slide 34: Common Errors
10. Lack of Training and Awareness
Error: Staff involved in distribution are unaware of their responsibilities under MDR Article 14, leading to oversight of critical checks.
Why it happens: Infrequent training or failure to update staff on regulatory changes.
Solution:
  Organize regular training sessions on MDR requirements.
  Develop a training matrix, identifying roles and required knowledge.
  Maintain training records to show compliance during audits, using eLearning for example.
  Periodically review, update and communicate your procedures.
By addressing these common errors with clear processes, training, and tools, distributors can better ensure compliance with MDR Article 14 and avoid regulatory issues.
Notes: Lastly, a common error that has been mentioned throughout this presentation is for there to be an overall lack of awareness or training in the requirements of the medical device regulations and in particular, the distributor obligations under article 14. Insufficient training can be because of lack of time or prioritization of the issues in a busy environment, but it is clear that compliance with these requirements should be considered important. Distributors should have appropriately trained staff, and should regularly review the regulatory situation and applicable legislation and keep records of training in these topics to demonstrate this compliance.
These errors all represent issues that distributors may encounter when attempting compliance with their regulatory obligations. These can generally be mitigated with an appreciation of these requirements through training, development and use of suitable procedures to handle these situations and implementing tools to help users record and process this information. Doing so will mean that issues can be handled more effectively and quickly, regulatory compliance is maintained and customers receive products that meet these requirements.

## Slide 35: Conclusion and Next Steps
Key Takeaways:
Distributors play a critical role in MDR compliance.
Collaboration ensures patient safety and regulatory alignment.
Next Steps:
Apply the knowledge in daily operations.
Share feedback and suggestions for continuous improvement.
Notes: To summarize this presentation, we want to be clear that under the European medical device regulation that distributors now play a key role in ensuring that devices on the market are compliant with legal requirements. Although the MDR is a complex document with a range of formal requirements, in practice many of these are generally common sense when it comes to confirming that products are as you expect, are kept in suitable environmental conditions and that you communicate effectively with other parts of the supply chain and with the appropriate legal authorities.
This collaboration with the manufacturer, customers and regulators helps ensure that the devices reach users and patients in a safe condition, with the right information and the correct legal documentation.

## Slide 36
Thank You!
For further questions or support, contact SP Medical A/S, Møllevej 1, DK-4653 Karise:
Philip James Davenport, pjd@sp-medical.dk, or
Hildur Annika Jacobsen, haj@sp-medical.dk